## Slide 1: Sigma Matrix and 6 Sigma
Dr. Bubul Kalita
Post Graduate Resident
Department of Biochemistry
Gandhi Medical College, Bhopal

## Slide 2: What is Six sigma?
Six sigma is a statistical approach to reduce variations (defects) in order to “Increase effectiveness”.
Sigma (σ) – Standard Deviation
Six sigma = 99.99967% effectiveness

## Slide 3
-1
6
5
4
2
3
1
0
50% Defect rate
31.6%
3.4 DPMO
More than 50% defect rate <<< sigma will be negative
If defect reduces>>>>>>>> Sigma level increases

## Slide 4: Lean
Lean is identification and elimination of waste from the process to “Increase Efficiency” without compromizing on the quality.
Example: Biomedical waste management

## Slide 5: History of Six Sigma
TOYOTA MOTORS
MOTOROLA >>>>> 1986
CEO: Bob Galvin
Bill Smith
(Father of Six Sigma)
Dr. Mikel Harry
Dr. Mikel Harry started the first six sigma academy in 1996

## Slide 6
General Electric (GE): 1995
He implemented 6 sigma in 5 different stages:
D (Define) : Selection of the right problem.
M (Measure): Collecting reliable data.
A (Analyze): To analyze Statistically to find the root cause.
I (Improve) : strategies to improve the problrm.
C (Control): Sustain the improvement.
CEO: Jack Welch
D M A I C

## Slide 7: How do we calculate Sigma metric?
Sigma =
Bias = Inaccuracy (Difference from reference value or peer group)
CV (coefficient of variation) = Imprecision/ reproducibility
TEa = Total Allowable Error (or “Tolerance”) required for that analyte. According to CLIA (Clinical Laboratory Improvement Amendment)
TEa ̶ ׀Bias׀
CV

## Slide 9
For Bias,
Individual Deviations: ∣Measured Value−True Value ∣
Average Deviation:
Express Bias as a Percentage:
Bias (%)= ×100
∑ Individual Deviations​
Number of Measurements
Average Deviation
True Value
For CV,
CV = ×100
Standard Deviation (SD)
Mean
For TEa
The Clinical Laboratory Improvement Amendments of 1988 (CLIA) Guidelines has given TEa of 85 analytes as per the criticality of the analyte.

## Slide 10: Method Decision chart
OPSpec Chart

## Slide 11: Assessing a single analyte across multiple system
The method decision chart is a tool used to assess a single analyte across multiple systems.
It helps in evaluating the performance of a single analyte across different testing methods and determining the best method for testing that analyte.

## Slide 12: Assessing multiple tests on one system across Labs
Taher et. al., 2018 .
Multi-site study (11 Hospitals in 9 countries) evaluating 18 tests on the Alinity Assay System

## Slide 13
1
2
3
4
5
1
5
2
3
4
10
6
7
9
8
Allowable Inaccuracy (% Bias)
Allowable Imprecision (%CV)
1σ
2σ
3σ
4σ
5σ
6σ
For Cholesterol,
TEa = 10%, Bias = 3%, CV = 3%
Sigma = (TEa% – Bias %)/ CV%
= (10-3)/3
= 7/3 = 2.33
Sigma = (TEa% – Bias %)/ CV%
CV% = (TEa% – Bias %)/ Sigma
Bias% = TEa% – (Sigma *CV%)
So for Bias = 0%,
CV% = TEa%/ Sigma
For CV% = 0,
Bias = TEa% - (Sigma* 0) = TEa = 10
For Sigma = 2, CV% = 10/2 = 5
For Sigma = 3, CV% = 10/3 = 3.33
For Sigma = 4, CV% = 10/4 = 2.5
For Sigma = 5, CV% = 10/5 = 2
For Sigma = 6, CV% = 10/2 = 1.66
*
2.33

## Slide 14: Normalized OPSpec Chart

## Slide 15: Levey-Jennings control chart

## Slide 16: Westgard Rules

## Slide 17: Evolution of Sigma Based QC
Original Westgard Rules, (March 1981)
Westgard JO et. al. Clin. Chem. Mar: 27(3); 493-501
Westgard Sigma Rules, (2014)
Westgard JO, Westgard SA,
Basic Quality Management System, 2014

## Slide 19: Westgard Sigma Rules with run size (2018)
Sigma metric | Control rule | QC frequency
6 | 13s, n=2 | 1 per 1000 patients samples
5 | 13s/22s/r4s, n=2 | 1 per 450 patients samples
4 | 13s/22s/r4s/41s, n=4 | 1 per 200 patients samples
3 | All “Westgard Rules”, n=6 | 1 per 45 patients samples

## Slide 20: Measurement Uncertainty(MU)
The International Vocabulary of Metrology (VIM) defines measurement uncertainty as a "non-negative parameter characterizing the dispersion of the quantity values being attributed to a measurand, based on the information used.“
Measurement uncertainty is often expressed as a Standard Deviation (SD) or Coefficient of variation (CV).

## Slide 21: Measurement uncertainty using Sigma Matrix
Sigma matrix = (TEa% - bias%)/ CV%
CV% = (TEa% - bias%)/ Sigma matrix
If want to provide world class quality, Sigma matrix = 6
CV% = (TEa% - bias%)/ 6
Sigma matrix 4.5 is for providing standard quality
CV% = (TEa% - bias%)/ 4.5
Example: For Cholesterol, TEa = 10, Bias = 3
CV% = (10-3)/6 = 1.166
CV% = (10-3)/4.5 = 1.55
In the clinical laboratory, there are some test parameters like Bilirubin, Sodium, Potassium; where maintaining CV% at very low level is very difficult. For such parameters SD is targeted.
SD = (TEa - bias)/ 4.5

## Slide 22: Impact on the laboratory
Sigma Level | DPMO | COPQ (Cost of Poor quality)
1 | 6,90,000 | > 40% of Revenue
2 | 3,08,537 | > 40% of Revenue
3 | 66,807 | 24 to 40% of Revenue
4 | 6,210 | 15 to 25% of Revenue
5 | 233 | 5 to 15% of Revenue
6 | 3.4 | < 1% of Revenue
Mahmood TS et. al. The Application of Six Sigma Methodology to Improve Service Quality: A Case Study in an Iraqi Retail Company. 2021 [cited 2024 Jan 15]
Poor quality:
More QC measurements and more rules to detect any deviation.
More delay, more repeat, more results.
Clinician and patient dissatisfaction.
Financial impact.

## Slide 23: Status in India
Nearly 40% of the hospitals surveyed did not apply any quality initiatives, and only 15 hospitals implemented Lean techniques, while 14 hospitals implemented Six Sigma techniques
The major reasons for not implementing Lean and Six Sigma initiatives in Indian healthcare are a lack of knowledge and availability of resources.
The implementation of Lean and Six Sigma in the Indian healthcare sector can lead to improved healthcare quality, reduced medical errors, and enhanced patient satisfaction.
(Suman and Prajapati. Utilization of Lean & Six Sigma quality initiatives in Indian healthcare sector. Rajiah K, editor. PLOS ONE. 2021 Dec 23;16(12):e0261747)

## Slide 24: Thank You